Automatic External Defibrillator (AED) Inspection Form

Location:

In premises where an AED is in place, checks must be carried out to ensure the device continues to be ready for use in the event of an emergency.

- Daily visual check: Check that the indicator light on the handle is GREEN.

If light is GREEN the device is rescue ready and no further action is needed.

Initial the table below to confirm that daily checks are carried out and record any faults where required.

|  | 1 | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Jan |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Feb |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| March |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| April |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| May |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| June |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| July |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| August |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| September |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| October |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| November |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| December |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| FAULTS: If light turns to red (with black cross) and/or the device sounds every 30 seconds there is a fault. If this occurs Lift the lid – the device will explain the nature of fault as follows: - Check pads: pads expired or disconnected - Battery low: new battery will be required approx. every 4 years - Maintenance/ service required: where this is indicated call Stewart Finlay on 01563 503456. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Nature of fault and action taken |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Date: |  |  |  |

- Monthly Check: Lift the lid and check the voice prompt is working. If you can hear the automated voice close the lid and no further action is needed.

Initial and date the table below to confirm monthly check has been carried out and record any faults where required.

| Monthly Check | Jan | Feb | March | April | May |  | June | July |  | August | Sept | Oct | Nov | Dec |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Date of check |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initial |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| If the voice prompt is not working call: Stewart Finlay on 01563 503456 for advice. |  |  |  |  |  | Date fault detected and action taken |  |  |  |  |  |  |  |  |

- Annual Check

Lift the lid, take the pads out and close the lid.

When the pads are removed, the alarm should sound and the indicator light should turn from green to red - this indicates that the device is working correctly and ready for use in the event of an emergency.

Reconnect the pads and close the lid. The indicator light should now return to green and no further action is required.

| Date of check | Annual Check By (Initial) | Working correctly Yes/ No | If the alarm does not sound or the light does not change colour when the pads are removed call Cardiac Science Technical Support on 0161 926000 and record the action required. |
| --- | --- | --- | --- |

- Further Information

- Pads should be replaced every two years. The date of expiry is visible on the front on the device.

Note: Newer G5 models will also warn when the pads are due to be replaced, by audible alarm

- Batteries will need to be replaced approximately every four years – both G3 and G5 units will warn when the battery is due to be replaced by audible alarm.

Replacement Pads and batteries are available from Stewart Finlay on 01563 503456